Responsibilities for Medical Locations

| Issue | Responsible Hospital Department | Reference Guidance Document | Notes |
| --- | --- | --- | --- |
| All Contactor Controls | Estates | 1. AXREM Installation Checklist for permanently wired medical devices https://www.axrem.org.uk/resource/electrical-installation-checklist/ 2. AXREM Guidance on contactor controls https://www.axrem.org.uk/resource/medical-device-installation-safety/ https://www.axrem.org.uk/resource/axrem-emergency-button-signs-for-permanently-installed-medical-equipment/ | Routine function testing may be carried out on behalf of Estates (EG RPAs) but records are held by Estates. The height should be 1.2 meters. |
| Door warning lights | RPA | AXREM Installation Checklist | RPAs need to be involved in the location decision on the Electrical Safety Group. |
| Disconnector Location | RPA | AXREM Installation Checklist | Disconnector refers to the means of electrical isolation. Contactors are used because large manual isolators may need some strength to operate in an emergency. The location of the contactor EPO control buttons is how IRR compliance is met. |
| Medical location diagrams documents | Estates | BS7671 A4 Requirement 710.514.9.101 | Estates should share these with colleagues who check medical radiological equipment |
| UPS load for permanently wired medical equipment | Relevant clinical team and or clinical engineering | AXREM Installation Checklist Note 4 | This is given in kVA |
| UPS Load for plug in medical devices | Clinical Engineering | AXREM Installation Checklist Note 4 | Total kW is acceptable. Clinical engineering should establish any future equipment that may be used. Medical Surgical Robot equipment can be quite a significant load. |
| UPS time | Clinical Staff | AXREM Installation Checklist Note 4 | The UPS Time should be provided by the senior clinical staff who undertake the clinical procedures. |
| Location and number of supplementary equipotential bonding connection points | Clinical Engineering | BS7671 A4 Requirement 710.415.2.101 | “The number of supplementary equipotential bonding connection points shall be determined by the designer and those responsible for ME equipment use and recorded as part of the information provided in Regulation 710.514.9.101”. This is a total change from the previous requirements. |
| Measurement of supplementary equipotential bonding | Estates Installing electricians | BS7671 A4 Annexe B710 710.514.9.101 (g) | The table is B710 (Informative) This means that people who test medical radiological equipment after installation only need to ask for this document. https://medical-locations.co.uk/about/final-testing SEB Measurements – typical example of a Record Schedule |